## Sheet1
S.No. | Product Name | Instrument Name | Test name | event Description | event ObservedIn | Description of Incident | Root Cause | Root Cause category.
1 | Milnacipran Hydrochloride (MI) | HPLC | Related Substances/Compounds by HPLC | Wrong weights entry | Finished product | During the review of Milnacipran Hydrochloride Related Compounds by HPLC, for known impurity (H-MCHRC01) weight wrongly entered in custom fields. (i.e.: H-MCHRC01 value entered as 2.125mg instead of 3.137mg). | Based on the above investigation, it is identified as analyst by inadvertently the impurity (H-MCHRC01) weight in-correctly entered in custom fields amounts table 2.125mg instead of 3.137mg. Hence, it was concluded as human error by the analyst. Reviewer also not ensured the abnormality while reviewing the sample set. | Human Error
2 | Levetiracetam (Form-I) (LE) | HPLC | Related Substances/Compounds by HPLC | Instrument Error / Failure observed | Finished product | During the analysis of Levetiracetam Related compounds by HPLC, while injecting Batch No.: LE23120536 test preparation-01, instrument failure observed. | Based on the investigation, obvious error observed as while injecting the 1st batch (LE23120536 Pre-01) injection, instrument failure occurred as Missing vial due to analyst by inadvertently sample was not prepared and not kept at respective vial position (Vial No.: 4). Hence, it is concluded as human error. | Human Error
3 | Terbinafine Hydrochloride (TN) | HPLC | Assay by HPLC | Peak(s) not eluted | Stability studies | During the analysis of Terbinafine Hydrochloride Assay by HPLC (On dried basis), it was noticed that principal peak (Terbinafine Hydrochloride) was not eluted in batch numbers [TN18050068 (LT) (60M) Pre-02 and TN18120157 (LT) (60M) Pre-02]. | Based on the investigation, initial analysis at sample solution injections principal peak (Terbinafine Hydrochloride) was not eluted due to injection volume was not taken during the injection due to needle was bent. Injection was injected on the vial cap instead of septa. It was happened due to misalignment of the sample tray. It was confirmed and rectified by the service engineer. | Instrument Error
4 | Ramipril (RMQ-RQ) | Particle size analyzer | Particle Size by Laser Diffraction | % RSD failure | Finished product | During the analysis of Ramipril Particle size by Laser diffraction (Wet method) for the batch RQ23040002, observed that %RSD not met the acceptance criteria at d(0.1). Observed RSD at d(0.1) is 15.4 % against acceptance criteria NMT: 15.0%. | Based on the above investigation, initial analysis %RSD not met the acceptance criteria due to analyst error. i.e. While preparation of RQ23040002 Pre-02 sample was not sonicated properly. Hence, peak pattern was not good when compared to regular trend and Pre-01. Due to this reason, %RSD not met the acceptance criteria. Hence, it is considered as human error. | Human Error
5 | Hydralazine Hydrochloride (HH) | HPLC | Content Method by HPLC | % RSD failure for bracketing standard solution | Finished product | During the analysis of Hydralazine Hydrochloride ‘Limit of Hydrazine content by HPLC’, it was observed that, in 1st bracketing Standard solution injection, % RSD for the areas of Hydrazine peak not met the acceptance criteria. Observed %RSD:6.69 %, Acceptance criteria: NMT: 2.0%. | Based on the above investigation, initial analysis at 1st bracketing standard solution injection ‘Hydrazine’ peak area was found lower side and %RSD not met the acceptance criteria might be due to improper purging of the injector/instrument prior to analysis. Hence, it was impacted at bracketing injection. Hence, it is considered as human error by the analyst. | Human Error
6 | Valsartan (VL) | HPLC | Related Substances/Compounds by HPLC | Vial number wrong | Finished product | During the Analysis of Valsartan Related compounds by HPLC(Method-II) in sample set (VL_FP_RC2_USP_270124_01), for 1st Bracketing System Suitability Solution injection vial number wrongly given as ‘2’ instead of ‘3’. | Based on the investigation, during the analysis at 1st bracketing system suitability solution injection, peak pattern not matched with the initial system suitability solution injection due to analyst in-correctly mentioned vail number as 2 instead of 3. Hence, standard solution injected in-place of 1st bracketing system suitability solution injection. Hence, it was concluded as human error by the analyst. | Human Error
7 | Terbinafine Hydrochloride (TN),Terbinafine Hydrochloride (TN),Terbinafine Hydrochloride (TN) | HPLC | Related Substances/Compounds by HPLC | Peak responses are not in line with the regular trend | Finished product | During the analysis of Terbinafine Hydrochloride Related Compounds Method-A by HPLC it is noticed that the bracketing system suitability solution peak pattern not matched with initial system suitability solution injection. | Based on the investigation, while running of analysis at 1st bracketing ‘system suitability solution’ injection, peak pattern was not matched with the initial ‘system suitability solution’ injection due to analyst by inadvertently updated sequence (i.e. vial number corrected sequence) was not injected into chromatographic system and injected initial created sequence [i.e. Which was in-correct vial number was mentioned (i.e. ‘3’ instead of ‘2’) for the 1st bracketing system suitability solution]. Due to this reason, sensitivity solution was injected in place of 1st bracketing system suitability solution. Hence, it was concluded as human error by the analyst. | Human Error
8 | NA | Autotitrator(Potentiometry) | Potentiometry | Data Invalid | Raw materials | During the analysis of RM211 Assay by Potentiometry, it is noticed that, after completion analysis result observed as invalid in report. | Based on investigation findings, the root cause for invalid result might be due to electrode lost sensitivity due to a layer of black coating formed around the silver ring of the electrode and analyst not ensured the electrode cleanliness prior to initiate the analysis. | Human Error
9 | Ezetimibe (Form-H) (EZ),Ezetimibe (Form-H) (EZ) | HPLC | Organic impurities by HPLC | RT shifting issue/ RT Not Matching with STP RT | Finished product | During the review of Ezetimibe Organic impurities Procedure-2 by HPLC, it is noticed that RT shifting observed from initial to bracketing system suitability solution and sensitivity solution injections. | Based on the above investigation, initial and further analysis RT shifting observed in sensitivity solution and system suitability solution injections due to outlet ball value buffers formed. It was confirmed by the external service engineer. Further, analysis instrument failure was occurred due to communication error. | Instrument Error,Communication Error
10 | Levetiracetam (Form-I) (LE) | HPLC | Assay by HPLC | Instrument Error / Failure observed | Finished product | During the analysis of Levetiracetam (Form-I) Assay by HPLC (On anhydrous and solvent-free basis) while running 5th batch of sample set B.No.: LE24010040 Inj-01, it is noticed that instrument failure occurred at around 0.5 minutes. | Based on the investigation, it was observed that, initial analysis while running of 5th batch (LE24010040 Pre-01), instrument failure occurred due to needle was bend due to sampling unit alignment missing. It was confirmed and rectified by the service engineer. | Instrument Error
11 | Levetiracetam (Form-I) (LE),Levetiracetam (Form-I) (LE),Levetiracetam (Form-I) (LE) | HPLC | Content Method by HPLC | Instrument Error / Failure observed | Stability studies | During the analysis of Levetiracetam (Form-I) Impurity-G by HPLC while running of 6th batch [LE21010011 (LT) (36M) Pre-01], it is noticed that instrument failure occurred at around 18 minutes and injection run time not completed as per STP run time as 25 minutes. | Based on the above investigation, in initial analysis while running of 6th batch [LE21010011 (LT) (36M) pre-01] instrument failure occurred at around 18 minutes, and injection run time not completed as per STP run time 25 minutes due to vacuum pump diaphragm (i.e. Inside the vacuum pump) dust formed and causing the degasser error. It was confirmed and rectified by the service engineer. | Instrument Error
12 | Hydralazine Hydrochloride (HH) | HPLC | Chromatographic purity by HPLC | Instrument Error / Failure observed | Intermediates | During the analysis of HHH-II (Intermediate of Hydralazine HCL) Chromatographic purity by HPLC, while running of 2nd batch (HHH2240046), it is noticed that instrument failure occurred at around 14 minutes and injection run time not completed, as per STP run time is 40 minutes. | Based on the above investigation, root cause identified as Communication error as loose connection of power cable to instrument identified. It might be happened due to analyst touched the power cables unknowingly and they might got disturbed. | Human Error
13 | Hydralazine Hydrochloride (HH) | HPLC | Organic impurities by HPLC | RT shifting issue/ RT Not Matching with STP RT | Finished product | During the analysis of Hydralazine Hydrochloride Organic impurities by HPLC, it is noticed that at 5th batch (HH24020043 Pre-02) Hydralazine Hydrochloride peak RT slightly varied when compared to remaining test solution injections in same sample set. | Based on the investigation, while running of analysis at 5th batch (HH24020043 Pre-02), Hydralazine Hydrochloride peak RT was slightly varied when compared to remaining test injections might be due to introduction of air bubble in to the flow path during the sample injection and causing the pressure drop. | Others
14 | Atomoxetine Hydrochloride (AM) | HPLC | Assay by HPLC | % RSD failure for bracketing standard solution | Finished product | During the analysis of Atomoxetine Hydrochloride Assay by HPLC’, it was observed that, in 1st bracketing Standard solution-02 injection, % RSD for the areas of Atomoxetine Hydrochloride peak not met the acceptance criteria. Observed %RSD:1.06 %, Acceptance criteria: NMT: 0.73%. | Based on the investigation, initial analysis at 1st bracketing standard solution-02 injection %RSD for the area of Atomoxetine Hydrochloride peak not met the acceptance criteria might be due to improper stabilization of instrument/ Column prior to analysis. | Human Error
15 | Darunavir Amorphous (DJ) | Particle size analyzer | Particle Size by Laser Diffraction | % RSD failure | In-Process samples | During the analysis of Darunavir Amorphous Particle size by Laser diffraction (Dry method) for the In-process batch DJ24020002, it was identified that %RSD not met the acceptance criteria at d (50) and d (90). Observed value: 16.8% at d (50) and 111% at d (90), acceptance criteria: NMT:10% and NMT:15% respectively. | Based on the investigation, initial analysis %RSD not met the acceptance criteria at d (50) and d (90) not met the acceptance criteria due to, in 2nd measurement abnormal peak (Ghost peak) was occurred. It was happened due to improper cleaning of sample flow line before the 2nd measurement. Hence, it is considered as human error by the analyst. | Human Error
16 | Zonisamide (ZM) | HPLC | Assay by HPLC | Instrument Error / Failure observed | Finished product | During the analysis of Zonisamide Assay by HPLC(on anhydrous basis) while running of 4th batch [ZM24020015 Pre-01], it is noticed that instrument failure occurred at around 10 minutes and injection run time not completed as per STP run time as 30 minutes. | Based on the investigation, initial analysis at 4th batch (ZM24020015 Pre-01) instrument failure occurred as Stop flow from remote host due to analyst by overlook pressed stop flow icon on the Empower screen. Hence, it was considered as human error by the analyst. | Human Error
17 | NA | Gas Chromatography | Chromatographic Purity by GC | Instrument Error / Failure observed | Recovery Solvents | During the analysis of Methylene dichloride (CODE:RM203) Distilled from DOL-II Chromatographic purity by GC, while running the sample solution injection having the batch number (DOL2RM203(D)240078) injection at around 22 minutes, instrument failure occurred and injection run not completed as per STP run time as 35 minutes. | Based on the above investigation it was observed that in initial analysis instrument failure was occurred while running of sample solution due to communication issue between instrument and computer it is concluded by Global Waters team. | Communication Error
18 | Dolutegravir Sodium (DG) | HPLC | Related Substances/Compounds by HPLC | Incomplete run or injection or measurement | Intermediates | During the analysis of DOL-II (Intermediate of Dolutegravir Sodium) Related compounds by HPLC, while running of system suitability solution bracketing injection of sample set DOL2_IN_RC_270224_01 it is noticed that, data incomplete occurred at about 2.0 minutes and injection run time not completed against run time 50.0 minutes. | Based on the above investigation, it was observed in system suitability solution bracketing injection incomplete run observed due to the kernel power manager has initiated a shutdown transition. It was happened due to Mr. P. Jaipal (Emp.ID: 212677) [who was performed on the instrument (QCD/265)] missed out to logout the system and another analyst Mr. J. Srinivas (Emp.ID: 132078) [analysis planning the instrument QCD/267)] inadvertently selected the instrument (QCD/265) monitor instead of instrument (QCD/267) monitor and already the system login analyst Mr. P. Jaipal (Emp.ID: 212677) without ensuring the another analyst Mr. J. Srinivas (Emp.ID: 132078) inadvertently restarted the computer causing the kernel power manager has initiated a shutdown transition of running analysis on the instrument (QCD/265. It was confirmed by IT personnel. Hence, it is concluded as human error. | Human Error
19 | Tenofovir Alafenamide Fumarate (TG) | HPLC | Enantiomeric Content/Purity by HPLC | Peak(s) not eluted | Finished product | During the analysis of Tenofovir Alafenamide Fumarate, ‘Diastereomeric Purity by HPLC it was observed that H-TAFRC03 peak is not eluted in bracketing System suitability solution injection. | Based on the investigation, initial analysis at ‘bracketing system suitability solution’ injection ‘H-TAFRC03’ peak not eluted due to the same peak was merged with the principal peak and further analysis the Resolution between ‘H-TAFRC03 and Tenofovir Alafenamide fumarate’ peaks not met the acceptance criteria due to column poor performance. | Column failure
20 | Terbinafine Hydrochloride (TN) | HPLC | Related Substances/Compounds by HPLC | Peak splitting or improper peak shape | Finished product | During the review of Terbinafine HCl Related Substances by HPLC (Method-A), it is noticed that in all injections Terbinafine HCl peak shape found not good (i.e. Fronting) was observed. | Based on the investigation, initial analysis in all injections ‘Terbinafine Hydrochloride’ peak shape was found not good (i.e. fronting observed) due to column poor performance. In further analysis the resolution between the ‘Impurity-B and Terbinafine Hydrochloride’ peaks not met the acceptance criteria due to insufficient column washing. | Column failure,Human Error
21 | Darunavir Amorphous (DJ) | KF Autotitrator | Water Content/ Moisture Content | Data Invalid | In-Process samples | During the Water content by KF analysis of DRS-I (solid) 5th interval for the B.No: DRS1240004 (5th Interval) Middle mixed Determination result found invalid. | Based on investigation findings, the root cause for invalid result in initial analysis is due to analyst [Mr. V.Kishore nana (Emp ID: 220505)] not entered the sample size details in application by inadvertently and started the analysis and in 2nd time analysis, analyst [Mr. C. Akshay (Emp ID: 223969)] selected wrong method i.e. ‘Water content (Above 1%)’ Method selected instead off “DRS-I” Method by inadvertently. Hence, it is concluded as a human errors. | Human Error
22 | Zonisamide (ZM) | DSC | Melting Range / Melting Point | Instrument Error / Failure observed | Finished product | While running the Zonisamide ‘Melting point by DSC’ for the Batch number ZM24020024, it was observed that a pop up displayed on Software window. | Based on the above investigation, root cause was identified as analyst RCS (Refrigerator cooling system) 90 was selected instead of RCS 40 due to wrong procedure mentioned in the SOP. | Human Error
23 | Ramipril (RP),Ramipril (RP),Ramipril (RP) | HPLC | Related Substances/Compounds by HPLC | Instrument Error / Failure observed | Finished product | During the analysis of Ramipril Related compounds by HPLC after completion blank solution which was injected before test sample B.No: RP24020008 it is noticed that instrument error occurred. | Based on the investigation, while running of blank solution injection (i.e. which was injected before test sample B. No.: RP24020008) instrument failure reported is due to system virtual memory exhausted issue and it was happened because of PC is not restarted from long back. It was confirmed by service engineer. | Communication Error
24 | Oseltamivir Phosphate (OV) | Particle size analyzer | Particle Size by Laser Diffraction | % RSD failure | Finished product | During the analysis of Oseltamivir Phosphate Particle size by Laser diffraction (Dry method) for the batch OV24020005, it was identified that % RSD not met the acceptance criteria at d (50) and d (90). Observed value: 18.9% at d (50) and 56.8% at d (90), acceptance criteria: NMT:10% and NMT:15% respectively. | Based on the investigation, initial analysis %RSD not met the acceptance criteria at d (50) and d (90) due to, in 2nd measurement abnormal peak (Ghost peak) was occurred due to improper cleaning of sample flow line before 2nd measurement. Hence, it was considered as human error by the analyst. | Human Error
25 | NA | Gas Chromatography | Residual solvents by GC | Peak(s) not eluted | Raw materials | During the analysis of RM048 Residual solvents by GC, it was observed that in all standard solution injection Dichloromethane (MDC) peak was not eluted whereas remaining solvents (Hexanes, Ethyl acetate, Methanol, Ethanol and Acetonitrile) peak eluted as per STP (HRM-286). | Based on the above investigation it was observed that in initial analysis Dichloromethane peak was not eluted in any of standard solution injections (i.e. initial six standard injections and bracketing standard injection) due to column poor performance. | Column failure
26 | Irbesartan(Form-A) (IE) | Ion Chromatography | Quantification by IC | Peak(s) not eluted | Method transfers | During the analysis of Irbesartan Azide Content by IC (Method transfer), it was noticed that, in Test (Blend solution, accuracy solution-01, accuracy solution-02 and accuracy solution-03), Azide peak was not eluted. | Based on the investigation, in initial analytical method transfer analysis, in ‘Test blend and Accuracy solution’ injections, Azide peak was not eluted due to Azide peak might be merged with the peak at RT around 3.5 minutes due to column poor performance. | Column failure
27 | Abacavir Sulfate (AB) | HPLC | Enantiomeric Content/Purity by HPLC | Instrument Error / Failure observed | Stability studies | During the analysis of Abacavir Sulfate Enantiomeric Purity by HPLC, 2nd bracketing system suitability injection instrument failure occurred around 5minutes against run time 45 minutes | Based on the above investigation, initial analysis instrument failure occurred at ‘2nd bracketing system suitability injection’ at around 5 minutes and run time not completed against the STP run time 45 minutes due to buffer overflow error occurred due to communication error. Hence, restarted the CPU and instrument. It was confirmed by the IT personnel. Further, analysis in initial system suitability solution injection all peak areas are found abnormal due to system suitability solution preparation error as analyst not ensured the complete dissolution of system suitability solution. Hence, it can be considered as human error. | Communication Error,Human Error
28 | Voriconazole (VO),Voriconazole (VO),Voriconazole (VO),Voriconazole (VO) | Ion Chromatography | Quantification by IC | % RSD failure for bracketing standard solution | Stability studies | During the analysis of Voriconazole Related compound-F by IC, it was noticed that, in 2nd Bracketing system suitability solution (B) %RSD not meet the acceptance criteria. Observed value is 31.21%, against acceptance criteria NMT:10.0% and next injections peaks not eluted. | Based on the investigation, in initial analysis, at 2nd Bracketing system suitability solution-B %RSD not met the acceptance criteria and further injections Related compound-F peak not eluted due to water flow to suppressor regeneration got interrupted and total conductivity got increased and sensitivity of detector got decreased. It was happened due to EO gas regulator not working. Further analysis initial Reference solution(C) inj-01, Reference solution(C) inj-02 and Reference solution(C) inj-03, Related compound-F peak not eluted due to pressure drop 1600 to 1000 psi due to minor leakage for detector cell in tubing. It was confirmed and rectified by the service engineer. | Others
29 | NA | HPLC | Related Substances/Compounds by HPLC | Peak splitting or improper peak shape | Raw materials | During the analysis of RM072 Related compounds by HPLC test, it is noticed that in initial test solution [i.e. RM072230001(Retest)] injection RM072 peak shape found not good. | Based on the investigation, initial analysis RM072 peak shape was found not good in test sample injection due to column poor performance. Hence, inactivated the column (HLL1/LCC/0395). 1st Reanalysis RM072 peak shape was found not good in test sample injection due to improper mixing of mobile phase. Hence, it is concluded as human error. 2nd Reanalysis RM072 peak shape was found not good in test sample injection due to insufficient column washing prior to analysis. Hence, it is concluded as human error. | Human Error,Column failure
30 | Opicapone (ON) | Gas Chromatography | Residual solvents by GC | Instrument Error / Failure observed | In-Process samples | During the analysis of OPN-II, Residual solvents by GC (Method-II), it was observed that while running of ‘OPN2240001(4th Interval) Middle’ injection instrument failure was observed around 5.0 min run time and injection run time (i.e. 30) was not completed as per STP (ONINP-002-00). | •Based on the above it was observed that in initial analysis, while running of ‘OPN2240001(4th Interval) Middle’ injection instrument failure was observed around 5.0 min run time against injection run time (i.e. 30 min) was not completed as per STP due transfer line was damaged (broken). It was confirmed and rectified by service engineer. | Instrument Error
31 | NA | HPLC | Enantiomeric Content/Purity by HPLC | Peak splitting or improper peak shape | Raw materials | During the analysis of RM330 Enantiomeric purity by HPLC test, it is noticed that RM330 (R-Isomer) peak shape found not good for the B.No.: RM330230002 (Re-test). | Based on the investigation, in initial reference solution and test solution [i.e. RM330230002 (Re-test)] injection RM330 (R-Isomer) peak shape found not good due to insufficient column washing. Hence, it is considered as a human error. | Human Error
32 | Dutasteride (DR) | KF Coulometer | Water Content/ Moisture Content | Data Invalid | In-Process samples | During the analysis of DSR-III (Solid) water content by coulometer, it is noticed that, after completion analysis result observed as ‘invalid’ in report. | Based on the investigation, while performing the analysis water content by Coulometer, found and data not saved on data base due to communication issue between PC and instrument due to RS232 cable was gone bad due to inside pin damaged which was connected to the instrument. Hence, it was considered as instrument error. | Instrument Error
33 | Levetiracetam (Form-I) (LE),Levetiracetam (Form-I) (LE) | HPLC | Enantiomeric Content/Purity by HPLC | Vial missing | Finished product | During the analysis of Levetiracetam (Form-I) Limit of Levetiracetam R-Enantiomer by HPLC, it was noticed that, while injecting the ‘Blank solution’ (which is injected after completion of 2nd bracketing standard solution) injection, instrument failure observed as Missing vial. | Based on the investigation, obvious error observed as while injecting the blank Solution (which is injected after completion of 2nd bracketing standard solution injection), instrument failure occurred as Missing vial due to analyst by inadvertently blank Solution vial was not placed at respective vial location (Vial No.: 29). Hence, it is concluded as human error. | Human Error
34 | NA | Autotitrator(Potentiometry) | Potentiometry | Data Invalid | Raw materials | During the analysis of RM435 Chloride content by Potentiometry, it is noticed that, result observed as invalid in report. | Based on investigation findings, the root cause for invalid result might be due to measuring interface found dusty. Hence, it is concluded as a instrument error. | Instrument Error
35 | NA | FTIR Spectrophotometer | Identification by IR | Instrument Error / Failure observed | Packing material | During the analysis of PM081 for the B. No: PM081240001 Identification by IR test, after scanned the sample spectrum, scan failed was observed. | Based on the above investigation, while scanning started the sample (PM081240001) spectrum, instrument was hanged due to the communication error due to scan pair mirror was not rotating properly due to the loose contact of their screw and system was not initialized on that time. It is confirmed and rectified by service personnel. | Communication Error
36 | Hydralazine Hydrochloride (HH) | HPLC | Content Method by HPLC | % RSD failure for bracketing standard solution | Finished product | During the analysis of Hydralazine Hydrochloride ‘Limit of Hydrazine content by HPLC’, it was observed that, in 2’nd bracketing Standard solution injection, % RSD for the areas of Hydrazine peak not met the acceptance criteria. Observed %RSD:6.26 %, Acceptance criteria: NMT: 2.0%. | Based on the investigation, initial analysis at 2nd bracketing standard solution injection ‘Hydrazine’ peak area was found higher side might be due to carryover from the column and it was impacted 2nd bracketing standard solution injection and leads to %RSD failure. | Others
37 | Etoricoxib (EX) | HPLC | Assay by HPLC | System suitability parameters not meeting the acceptance criteria | Finished product | During the analysis Etoricoxib Assay by HPLC (On dried basis), it was observed that, % RSD not met the acceptance criteria for Etoricoxib Standard solution-2 at 2nd Bracketing. Observed value is 2.31% against the acceptance criteria NMT: 2.0%. | Based on the above investigation, initial analysis at 2nd bracketing standard solution-02 injection %RSD for the areas of Etoricoxib peak not met the acceptance criteria might not be solution mixed properly. Hence, it was considered as human error. | Human Error
38 | Terbinafine Hydrochloride (TN) | HPLC | Related Substances/Compounds by HPLC | System suitability parameters not meeting the acceptance criteria | Finished product | During the analysis Terbinafine Hydrochloride Related substances by HPLC, it was observed that, 1st Bracketing standard solution % RSD for the areas of Terbinafine Hydrochloride peak not met the acceptance criteria. Observed value is 3.38% against the acceptance criteria NMT: 2.0%. | Based on the investigation, initial analysis at 1st bracketing standard solution injection ‘Terbinafine Hydrochloride’ peak area was found higher side and %RSD not met the acceptance criteria due to overtightened the vial cap and the solution filled above the neck. Hence, it might be cause for the area difference and it leads to %RSD failure at 1st bracketing injection. Hence, it is considered as human error by the analyst. | Human Error
39 | Abacavir Sulfate (AB) | HPLC | Enantiomeric Content/Purity by HPLC | Instrument Error / Failure observed | Finished product | During the analysis of Abacavir Sulfate Enantiomeric Purity by HPLC, 2nd bracketing Reference solution(a) injection instrument failure occurred around 15 minutes against run time 45 minutes. | Based on the above investigation, initial analysis instrument failure occurred at ‘2nd bracketing Reference solution (a) injection’ at around 15 minutes against the STP run time 45 minutes due to communication error i.e. momentary displacement of the CPU due to analyst inadvertently hit the CPU stand. | Human Error
40 | Zonisamide (ZM),Zonisamide (ZM),Zonisamide (ZM) | HPLC | Related Substances/Compounds by HPLC | % RSD failure for bracketing standard solution | Finished product | During the analysis Zonisamide Related substances by HPLC, it was observed that, 1st Bracketing standard solution % RSD for the areas of Zonisamide peak not met the acceptance criteria. Observed value is 28.52% against the acceptance criteria NMT: 10.0%. | Based on the above investigation, in initial analysis at bracketing Standard solution injection, the % RSD for the area of Zonisamide peak not met the acceptance criteria due to carry over observed due to needle seat was partially blockage. It was confirmed and rectified by the service engineer. | Instrument Error
41 | Desloratadine (DL) | KF Coulometer | Water Content/ Moisture Content | Data Invalid | Stability studies | During the analysis of Desloratadine water content by coulometer, it is noticed that, DL19090001(LT)48M Determination-02 result observed as ‘invalid’ in report. | Based on the investigation, while performing the analysis water content by Coulometer having the batch number DL19090001 (LT)(48M) Determination-02 invalid result found due to analyst by sample weight was not given inadvertently by in tiamo application default taken as ‘0’ instead of 0.25403gr. It was identified as human error. | Human Error
42 | Entacapone (EP) | HPLC | Assay by HPLC | Instrument Error / Failure observed | Stability studies | During the analysis of Entacapone Assay (On Dried Basis) by HPLC, It was observed that instrument failure occurred around 18minutes against run time 30 minutes for Batch number EP22020003(Z-IV)24M Pre-01. | Based on the above investigation, initial analysis while running of 6th batch [EP22020003 (Z-IV) (24M) Pre-01] instrument failure occurred due to analyst by inadvertently opened the column compartment door while running the analysis. Hence, it was considered as human error. | Human Error
43 | Olmesartan Medoxomil (OQ) | HPLC | Related Substances/Compounds by HPLC | System suitability parameters not meeting the acceptance criteria | Stability studies | During the analysis of Olmesartan medoxomil Related compound by HPLC, it was observed that, initial to 2nd bracketing standard system suitability solution injections resolution between Olmesartan medoxomil and Related compound-A not met the acceptance criteria. | Based on the above investigation, initial analysis the Resolution between the Olmesartan Medoxomil and Related compound-A peaks not met the acceptance criteria due to insufficient column washing. Hence, it is considered as human error. | Human Error
44 | Hydralazine Hydrochloride (HH) | HPLC | Chromatographic purity by HPLC | Data file not saved after completion of analysis | Intermediates | HHH-II (Intermediate of Hydralazine HCL), Chromatographic purity by HPLC, it was observed that, after completion of the analysis, verified the data, data not generated for test solution injections. | Based on the above investigation, in initial analysis data not generated for test solution injections due to communication error between CPU and system due to SAT cable of CPU gone bad. It was confirmed by IT person and same was rectified. | Instrument Error,Communication Error
45 | NA | FTIR Spectrophotometer | Identification by IR | Data file not saved after completion of analysis | Raw materials | During the analysis of RM184 Identification by IR test, after scanned the test preparation (i.e: RM184240003 Container No:03), system was hanged and scan was not completed and data not saved in database | Based on the above investigation, while scanning started the sample (RM184240003 container-03) spectrum, instrument was hanged due to scan pair mirror stopped rotating and software gets hanged. It was confirmed by service personnel. | Instrument Error
46 | Voriconazole (VO),Voriconazole (VO),Voriconazole (VO) | HPLC | Related Substances/Compounds by HPLC | Peak(s) not eluted | Finished product | During the analysis of Voriconazole Related compounds by HPLC, while processing the chromatograms, it is noticed that “ EP Impurity-B” peak was not eluted in initial and bracketing standard solution injections. | Based on the above investigation, this laboratory incident occurred due to analyst incorrect impurity standard {i.e: USPRC-B impurity standard [i.e. VORRCD (Voriconazole Enantiomer), having the Lot No.: VOR/RC-D/01/22] used instead of ‘Related compound-D (EP Impurity-B)’ (Lot No.: VO002/17001)’ during the standard solution preparation. Hence, it is considered as human error. | Human Error
47 | Terbinafine Hydrochloride (TN) | HPLC | Chromatographic purity by HPLC | Peak(s) not eluted | Intermediates | During the analysis of TNF-I (Intermediate of Terbinafine HCL) Chromatographic purity by HPLC, peaks are not eluted in test solution (TNF1240072) injection. | Based on the above investigation, peaks not eluted in test solution injection due to analyst filled the blank solution in vial instead of test solution and placed the same in HPLC vial rack. Hence, it is considered as human error. | Human Error
48 | Oseltamivir (OM) | HPLC | Organic impurities by HPLC | Instrument Error / Failure observed | Stability studies | During the analysis of Oseltamivir Organic impurities by HPLC, in test sample B.No.:OM23030001(5°C)12M preparation-02 injection instrument failure. | Based on the investigation, during the analysis while injecting the ‘OM23030001(5°C) 12M Prep’n-02’ injection instrument failure occurred due to gripper struck and motor over temperature due to transport assembly not moving properly. It was confirmed by service engineer. | Instrument Error
49 | Oseltamivir Phosphate (OS) | HPLC | Content Method by HPLC | RT shifting issue/ RT Not Matching with STP RT | Stability studies | During the analysis of Oseltamivir Phosphate Triphenyl Phosphine Oxide Content by HPLC, it was observed that, ‘Triphenyl Phosphine Oxide (H-OLPRC05)’ peak RT variation observed between initial and Bracketing Standard Solution injection. | Based on the investigation, initial analysis in initial ‘standard solution’ injection, H-OLPRC05 peak RT was varied when compared to the bracketing standard solution injection due to air bubble might be formed in flow path during the analysis. Hence. RT was shifted. | Others
50 | Valsartan (VL) | HPLC | Related Substances/Compounds by HPLC | Peak(s) not eluted | Stability studies | During the analysis of Valsartan Related compounds by HPLC (Method-1), while processing the chromatograms, it is noticed that “H-VSNRC01 impurity peak not eluted in initial and bracketing reference solution injections. | •Based on the above investigation, initial analysis in initial and bracketing Reference solution injections ‘H-VSNRC01’ peak not eluted might be due to analyst by inadvertently VQL-I impurity used instead of H-VSNRC01. Hence, it was considered as human error. •In further analysis, in initial and bracketing Reference solution(b) injections ‘Related compound-B (VSRCB)’ peak area not regular trend, might be due to analyst by inadvertently ‘Valsartan’ Working standard used instead of Peak ID. Hence, it was considered as human error. | Human Error
51 | Eplerenone (PR) | HPLC | Related Substances/Compounds by HPLC | Instrument Error / Failure observed | Stability studies | During the analysis of Eplerenone Related compounds by HPLC while injecting the 4th batch [PR22110030(Z-IV)18M Pre-01] in sample set, it is noticed that instrument failure occurred. | Based on the investigation, while injecting the 4th batch [PR22110030(Z-IV)18M Pre-01] in sample set, instrument failure occurred as Missing vial due to analyst by inadvertently not prepared the sample at the time of injection acquisition and not placed the sample vial at respective vial location (Vial No.: 11). Hence, it is concluded as human error. | Human Error
52 | Irbesartan(Form-A) (IE) | HPLC | Related Substances/Compounds by HPLC | Peak(s) not eluted | Finished product | During the analysis of Irbesartan Related compounds (Method-A) by HPLC, in test solution for the B.No: IE24050049 Pre-02 Peaks are not eluted. | Based on the investigation, initial analysis in 1st batch (IE24050049 pre-02) peaks are not eluted due to leak observed at six-port valve due to SS-tubing from six port valve to column compartment got broken causing leakage. It was confirmed and rectified by service engineer. | Instrument Error
53 | Losartan Potassium (Form-I) (LR) | HPLC | Related Substances/Compounds by HPLC | Instrument Error / Failure observed | Stability studies | During the analysis of Losartan Potassium Related compounds by HPLC while injecting the Blank solution i.e. after completion of Bracketing Reference solution(a) in sample set, it is noticed that instrument failure occurred. | Based on the above investigation, instrument failure occurred due to sampler cooler area water condensed and accumulated and it was happened due to sampler cooler temperature is 5°C and the water has to be drained through drain tube whereas drain tube was blocked and observed that tube was Quenched. Hence, water came back in to cooler compartment and floated in to pump module, it leads to the instrument failure. It was confirmed and rectified by service personnel. | Instrument Error
54 | NA | HPLC | Related Substances/Compounds by HPLC | Peak responses are not in line with the regular trend | Raw materials | During the review of RM422 Related Compounds by HPLC, it was noticed that, in system suitability and test solution injections RM422 peak response not in regular trend. | Based on the investigation, initial analysis in ‘System suitability solution and test solution’ injections, RM422 peak area not in-line with the regular trend and AHS N-Acetyl impurity peak area not within the regular trend in system suitability injections due to instrument error i.e: rotor seal got blocked due to buffer precipitations deposited. | Instrument Error
55 | Aripiprazole (AO) | HPLC | Chromatographic purity by HPLC | Instrument Error / Failure observed | Intermediates | During the analysis of ARZ-II Chromatographic purity by HPLC, while injecting test Solution, instrument failure observed. | Based on the investigation, while injecting the batch ARZ2240010 test solution in sample set, instrument failure occurred as Missing vial due to analyst by inadvertently not prepared the sample at the time of injection acquisition and not placed the sample vial at respective vial location (Vial No.: 3). Hence, it is concluded as human error. | Human Error
56 | NA | HPLC | Related Substances/Compounds by HPLC | Peak(s) not eluted | Raw materials | During the analysis of RM239 Related compounds by HPLC, in test solution for the B. No: RM239230002 (Retest) Peaks are not eluted. | Based on the investigation, in test solution for the B. No: RM239230002 (Retest), peaks are not eluted due to outlet ball valve blocked. It was confirmed and rectified by the service engineer. | Instrument Error
57 | Irbesartan(Form-A) (IE) | Particle size analyzer | Particle Size by Laser Diffraction | % RSD failure | Finished product | During the analysis of Irbesartan Particle size by Laser diffraction (Dry method) for the batch IE24060059, it was observed that %RSD not met the acceptance criteria. Observed %RSD at d(0.9) is 125% against acceptance criteria NMT: 30.0%. | Based on the investigation, initial analysis %RSD not met the acceptance criteria at d (0.9) due to, in 2nd measurement abnormal peak (Ghost peak) was occurred due to dust collection bag was filled with powder and it might be impacted the vacuum and leads to ghost peak. | Instrument Error
58 | Aripiprazole (AO) | HPLC | Assay by HPLC | System suitability parameters not meeting the acceptance criteria | Finished product | During the analysis of Aripiprazole Assay by HPLC’, it was observed that, in bracketing Standard solution-02 injection, % RSD for the areas of Aripiprazole peak not met the acceptance criteria. Observed %RSD:1.14 %, Acceptance criteria: NMT: 1.0%. | Based on the above investigation, initial analysis in bracketing standard solution-02 injection %RSD for the areas of Aripiprazole peak not met the acceptance criteria due to bracketing standard solution-02 injection area observed higher side. It was happened due to sample precipitation observed around the injector tip and this precipitation might have carried to the bracketing standard injection through external surface of the tip. It was confirmed by external service engineer. | Instrument Error
59 | Losartan Potassium (Form-I) (LR),Losartan Potassium (Form-I) (LR),Losartan Potassium (Form-I) (LR),Losartan Potassium (Form-I) (LR) | HPLC | Acetic Acid content by HPLC | Instrument Error / Failure observed | Finished product | During the analysis of Losartan Potassium Acetic acid content by HPLC, while running the B.No:LPR24060020 pre-01, it was observed that instrument failure around 0.30 min and run not completed as per STP runtime (35 min). | Based on the above investigation, initial and further analysis instrument failure was observed at 1st batch (LR24060020 pre-01) due to over pressure occurred due to injector needle got blocked causing system over pressure. In further analysis, instrument failure was observed at 1st batch (LR24060020 pre-02) due to over pressure occurred at 0.15 minutes due to folding in the peak tubing at column inlet. | Instrument Error
60 | Dolutegravir Sodium (DU) | HPLC | Content Method by HPLC | RT shifting issue/ RT Not Matching with STP RT | In-Process samples | During the analysis of DUS-I (R. MASS-I) (In-process) RM339 content by HPLC, in ‘test solution’ injection of 3rd interval, the RT of DUS peak RT not matched with STP RT. | Based on the investigation, during the analysis in 3rd interval sample solution injection, ‘DUS’ peak RT was varied and not matched with the STP RT due to air bubble might be formed in flow path during the analysis. | Instrument Error
61 | Dolutegravir Sodium (DG) | Particle size analyzer | Particle Size by Laser Diffraction | % RSD failure | Finished product | During the analysis of Dolutegravir Sodium Particle size by Laser diffraction (Wet method) for the batch DG24060159 Preparation-01, in measurement-01 abnormal peak observed. | Based on the investigation, in initial analysis B. No.: DG24060159 Pre-01 at measurement-1 abnormal peak was observed due to tiny air bubbles might be present in the dispersant and which might be impacted the analysis causing ghost peaks. It was confirmed by the service engineer. | Others
62 | Vortioxetine Hydrobromide (VH) | KF Autotitrator | Water Content/ Moisture Content | Data Invalid | In-Process samples | During the Water content by KF analysis of VXH (solid) In-process sample for the B.No: VH24060021 Determination result found invalid. | Based on investigation findings, the root cause for invalid result in water content by KF analysis due to analyst [Mr. S. Ashok (Emp ID: 226101)] by inadvertently not entered the sample size details in tiamo application. Hence, it is concluded as a human error. | Human Error
63 | Irbesartan(Form-A) (IE) | Gas Chromatography | Residual solvents by GC | Instrument Error / Failure observed | Finished product | During the analysis of Irbesartan (Form-A), Residual solvents by GC (Method-II), instrument failure was observed while running of IE24060066 Prep-01 at RT about 11.5 minute and injection run time (25 min) not completed as per STP. | Based on the investigation, in initial analysis while running of 2nd batch (IE24060066 Prep-01) instrument failure was observed due to ‘Network link disconnected’ in middle of the analysis. It might be happened due to network LAN cable problem. Hence, same was replaced by service engineer. | Communication Error
64 | Entacapone (EP),Entacapone (EP),Entacapone (EP),Entacapone (EP),Entacapone (EP) | HPLC | Organic impurities by HPLC | % RSD failure for bracketing standard solution | Finished product | During the analysis of Entacapone Organic impurities by HPLC’, it was observed that, in 2nd bracketing System suitability solution injection, % RSD for the areas of Entacapone peak not met the acceptance criteria. Observed %RSD:17.29 %, Acceptance criteria: NMT: 10.0%. | Based on the investigation, in initial analysis at 2nd bracketing system suitability solution, %RSD for the areas of Entacapone peak not met the acceptance criteria due to septa gone bad. Hence, there might be chance to get evaporation of solvent mixture from vial and it leads to increasing the main peak area of 2nd bracketing system suitability solution injection. | Others
65 | Vortioxetine Hydrobromide (VH) | Particle size analyzer | Particle Size by Laser Diffraction | % RSD failure | Finished product | During the analysis of Vortioxetine Hydrobromide Particle size by Laser diffraction (Dry method) for the batch VH22120017, it was observed that, at d(0.9) %RSD not met the acceptance criteria. Observed %RSD at d(0.9) is 48.1% against acceptance criteria: NMT: 30%. | Based on the investigation, in initial analysis %RSD not met the acceptance criteria at d (0.9) due to abnormal peak (Ghost peak) was observed it might be due to improper cleaning of sample flow line before 2nd and 3rd measurement. Hence, it was considered as human error. | Human Error
66 | Dolutegravir Sodium (DU) | HPLC | Content Method by HPLC | Instrument Error / Failure observed | In-Process samples | During the analysis of DUS-I (R. MASS-II) (In-process) Dolutegravir Stage-IA Content by HPLC, in ‘Sample solution’ injection [i.e.: DUS1240012 (2nd interval)] instrument failure occurred around 14 minutes. | Based on the investigation, while running of B. No: DUS1240012 (2nd interval) sample solution injection, Instrument failure occurred due to intermitted disconnection of 8 port LAN switch and leads to connectivity lost between the instrument and computer. | Instrument Error
67 | Hydralazine Hydrochloride (HH) | HPLC | Content Method by HPLC | % RSD failure for bracketing standard solution | Finished product | During the analysis of Hydralazine Hydrochloride ‘Limit of Hydrazine content by HPLC’, it was observed that, in 1st bracketing Standard solution injection, % RSD for the areas of Hydrazine peak not met the acceptance criteria. Observed %RSD:3.32 %, Acceptance criteria: NMT: 2.0%. | Based on the investigation, initial analysis at 1st bracketing standard solution injection, %RSD for the area of ‘Hydrazine’ peak not met the acceptance criteria due to air bubble might be trapped into the injector and impacted the bracketing standard solution injection and leads to the area variation. | Others
68 | Irbesartan(Form-A) (IE) | Particle size analyzer | Particle Size by Laser Diffraction | % RSD failure | Finished product | During the analysis of Irbesartan (Form-A) Particle size by Laser diffraction (Dry method) for the batch IE24060069, it was observed that %RSD at d (0.9) not met the acceptance criteria. Observed %RSD at d (0.9) is 33.6% against acceptance criteria NMT: 30.0%. | Based on the investigation, initial analysis %RSD not met the acceptance criteria at d (0.9) due to, in 3rd measurement abnormal peak (Ghost peak) due to water accumulated in air compressor tank and might be moisture was formed and this moisture was entered in sample flow line it leads to ghost peak. It was happened due to non-return valve chocked. | Instrument Error
69 | Dolutegravir Sodium (DU) | Particle size analyzer | Particle Size by Laser Diffraction | % RSD failure | In-Process samples | During the analysis of Dolutegravir Sodium Particle size by Laser diffraction (Wet method) for the batch DU24070011 (1st Interval) Preparation-01, it was observed that in measurement -01 %RSD at d (90) not met the acceptance criteria. Observed %RSD at d (90) is 165% against acceptance criteria NMT: 30.0%. | Based on the investigation, in initial analysis B.No.: DU24070011 (1st interval) pre-01 at measurement-1 abnormal peak was observed due to analyst missed to mix the sample by using the cyclomixer for 2 minutes as per STP and transferred in to Hydro MV drop by drop and it might be impacted and leads to abnormal peak. | Human Error
70 | Ezetimibe (Form-H) (EZ) | HPLC | Organic impurities by HPLC | RT shifting issue/ RT Not Matching with STP RT | Finished product | During the review of Ezetimibe Organic impurities Procedure-2 by HPLC, it is noticed that RT’s are shifting observed from initial system suitability injection, sensitivity solution injections and test sample injections. | Based on the above investigation, in initial and further analysis Ezetimibe peak RT found inconsistence due pressure fluctuations. It was happened due to AIV cartridge got blocked. It was confirmed by external service engineer. | Instrument Error
71 | Irbesartan Intermediate ((1-[(2`-Cyanobiphenyl-4-yl)methyl]-2-n-butyl-4-spirocyclopentane-2-imidazolin-5-one) (IE) | HPLC | Assay by HPLC | System suitability parameters not meeting the acceptance criteria | Intermediates | During the analysis of IES (Intermediate of Irbesartan) Hodtime study of Assay by HPLC (On dried basis) test, it was observed that, in bracketing Standard solution injection, % RSD for the areas of IES peak not met the acceptance criteria. Observed %RSD: 2.34 %, Acceptance criteria: NMT: 2.0%. | Based on the investigation, in initial analysis at bracketing standard solution, %RSD for the areas of IES peak not met the acceptance criteria due to septa slit was opened slightly bigger when compared with other vial septas. Hence, solvent (Diluent is methanol) get evaporated (Diluent is methanol) from vial and it leads to increasing the main peak area of bracketing standard solution injection. | Others
72 | Hydralazine Hydrochloride (HH) | HPLC | Content Method by HPLC | % RSD failure for bracketing standard solution | Finished product | During the analysis of Hydralazine Hydrochloride ‘Limit of Hydrazine content by HPLC’, it was observed that, in 1st bracketing Standard solution injection, % RSD for the areas of Hydrazine peak not met the acceptance criteria. Observed %RSD:20.83%, Acceptance criteria: NMT: 2.0%. | Based on the investigation, initial analysis at 1st bracketing standard solution injection %RSD for the area of ‘Hydrazine’ peak not met the acceptance criteria due to area found lower side when compared to the remaining standard solution injections. It was happened due to minute leakage at the measuring pump due to measuring seal issue. | Instrument Error
73 | Hydralazine Hydrochloride (HH) | HPLC | Chromatographic purity by HPLC | Instrument Error / Failure observed | Intermediates | During the analysis of HHH-I (Intermediate of Hydralazine Hydrochloride) Chromatographic purity by HPLC, instrument failure was observed in 2nd test solution (i.e. HHH1240194) injection about 29 minutes and run time not completed as per STP (i.e. 50 minutes). Error observed as ‘Instrument failure AgilentLC#QCD_266’. | Based on the above investigation, root cause identified for Instrument failure as leak detected due to minor buffer precipitations on rotor seal causing blockage of rotor seal and it leads to back pressure causing leakage from 4th port of 6-port valve. Minor buffer precipitations occurred due to inadequate washing of the system prior to analysis by the analyst leads to blockage of rotor seal. | Instrument Error,Human Error
74 | Zoledronic Acid Monohydrate (ZL) | Gas Chromatography | Residual solvents by GC | % RSD failure for bracketing standard solution | Method transfers | During the analysis of Zoledronic acid monohydrate for Residual Solvents by GC (Method transfer) in 3rd bracketing Standard solution injections chlorobenzene %RSD not met the specification limit. Observed % RSD 11.44 against Acceptance criteria NMT: 10.0%. | Based on the investigation, during the initial analysis it was observed that in the 3rd bracketing (which was injected for the batch analysis and Accuracy for spikes solution) standard solution injections chlorobenzene % RSD not met the acceptance criteria due to the improper crimping of the HSS vial. Based on the investigation, during the reanalysis it was observed that the % recovery for MDC and Chlorobenzene solvents in spiked/recovery solutions are not within the acceptance criteria due to the cyberlab vials. | Human Error
75 | Atomoxetine Hydrochloride (AM) | HPLC | Enantiomeric Content/Purity by HPLC | Peak(s) not eluted | Stability studies | During the analysis of Atomoxetine HCL Isomeric Purity by HPLC, it is noticed that Atomoxetine HCL peak was not eluted second Bracketing Reference solution (b) injection. However, remaining impurities (Impurity-B and Impurity-D) peaks are eluted. | Based on the investigation, initial analysis in 2nd bracketing Reference solution(b) injection ‘Atomoxetine Hydrochloride’ peak not eluted due to air bubble might be trapped and interrupted in to flow line path and it leads to the air lock and impacted at that particular time of injection. | Instrument Error
76 | Levetiracetam (Form-I) (LE) | HPLC | Content Method by HPLC | Peak(s) not eluted | Finished product | During the analysis of Levetiracetam(Form-I) Acetic acid content by HPLC, it is noticed that Acetic acid peak was not eluted second Bracketing standard solution injection. | Based on the investigation, in initial analysis ‘Acetic acid’ peak was not eluted in 2nd bracketing standard solution injection due to pressure was dropped at the time of ‘Blank Solution’ injection and it was continued to the 2nd bracketing standard solution injection. It was happened due to check valve cartridge partially blocked. It was confirmed and rectified by service engineer. | Instrument Error
77 | Lopinavir (LQ) | FTIR Spectrophotometer | Identification by IR | Data file not saved after completion of analysis | Finished product | During the analysis of Lopinavir Identification by IR test, after scanned the test sample having the B. No.: LQ24070005, it is noticed that, Test data not saved on data base. | Based on the above investigation, while scanning started the sample (LQ24070005) spectrum, instrument was rebooted. It was happened due to system not responding. Hence, data not generated in data base. It was identified and confirmed by service personnel. | Instrument Error
78 | Oseltamivir Phosphate (OV) | X-RAY DIFFRACTOMETER | Identification by XRD | Others | Stability studies | During the execution of Oseltamivir phosphate XRD Identification test for B. No.: [OV24050052(ACC)(1M)], error displayed as MS_SC09 Tried to load or preload a sample from position “1F06” position was empty. | Based on the investigation, while scanning of ‘B. No.: [OV24050052(ACC)(1M)]’ error occurred due to while creating the job creation analyst by inadvertently for the particular sample position was in-correctly selected as “1F06” position instead of “1A06” position. Hence, it leads to the error as ‘MS_SC09 Tried to load or preload a sample from position “1F06” position was empty’. Further, verified the practice and procedure and observed that once the job created, the same was not verified the reviewer. | Human Error
79 | Ezetimibe (Form-B) (EQ) | HPLC | Organic impurities by HPLC | Peak responses are not in line with the regular trend | Finished product | During the analysis of Ezetimibe (Form-B) Organic impurities (Procedure-I) by HPLC, it was noticed that, in initial system suitability solution injection, Ezetimibe and O-fluorobenzene isomer peaks response not in-line with regular trend. | Based on the investigation, initial analysis in initial system suitability solution injection Ezetimibe’ and ‘O-fluorobenzene isomer’ peak areas found not in-line with regular trend due to less volume was observed in vial. Hence, suspecting that, air might be entered in to the injector and air gap formed, due to this, less volume injected in to the system and impacted the areas. | Human Error
80 | Losartan Potassium (Form-I) (LR) | HPLC | Acetic Acid content by HPLC | Instrument Error / Failure observed | Finished product | During the analysis of Losartan Potassium Acetic acid content by HPLC while running 2nd batch of sample set B.No.: LR24070032 Preparation-02, it is noticed that instrument failure occurred at around 0.16 minutes. | Based on the investigation, initial analysis instrument failure was observed at 2nd batch (B. No.: LR24070032 Preparation-02) might be occurred due to obstruction of minute particle in peak tubing which is connected from the HPLC injector port to column inlet, leads to blockage and causing over pressure. | Instrument Error
81 | Dolutegravir Sodium (DU) | HPLC | Content Method by HPLC | Peak(s) not eluted | In-Process samples | During the analysis of DUS-I (In-process sample) RM339 Content by HPLC, in ‘Standard and test solution’ injections peaks are not eluted. | Based on the investigation, during the initial analysis peaks not eluted in standard and test solution injections might be due to leakage at injector syringe causes to injection volume was not taken properly. | Instrument Error
82 | Levetiracetam (Form-I) (LE) | Particle size analyzer | Particle Size by Laser Diffraction | Others | Finished product | During the analysis of Levetiracetam (Form-I) Particle size by Laser diffraction (Dry method) for the batch LE24040110 (Customer sample-02), it is observed that, at Measurement-1 abnormal (Ghost) peak was identified. | Based on the investigation, initial analysis for the batch LE24040110 (Customer sample-02) 1st measurement abnormal peak (Ghost peak) was identified might be due to improper cleaning of the sample flowline. | Human Error
83 | Milnacipran Hydrochloride (MI) | HPLC | Related Substances/Compounds by HPLC | Peak responses are not in line with the regular trend | Stability studies | During the review of Milnacipran HCl Related Compounds by HPLC analysis data, it was noticed that in ‘sensitivity solution’ injection “Milnacipran HCl” peak area observed high compared to regular trend. | Based on the investigation, initial analysis in sensitivity solution injection Milnacipran hydrochloride peak area found not in-line with the trend might be due to sensitivity solution dilution error. Further analysis in initial Reference solution injection H-MCHRC01 peak was not eluted due to analyst H-MCHRC01 standard might not be transferred in to volumetric flask. | Human Error
84 | Levetiracetam (Form-I) (LE) | Particle size analyzer | Particle Size by Laser Diffraction | % RSD failure | In-Process samples | During the analysis of Levetiracetam (Form-I) Particle size by Laser diffraction (Wet method) for the batch LE24080293-INP, it is observed that, %RSD not met the acceptance criteria at d(0.9). Observed value:21.7%, Acceptance criteria NMT:15.0%. | Based on the investigation, in initial analysis %RSD not met the acceptance criteria at d(0.9) and agglomerates occurred in both preparations might be due to analyst after adding the dispersion medium stirring was not performed properly to form uniform dispersion by inadvertently. | Human Error
85 | Aripiprazole (AO),Aripiprazole (AO),Aripiprazole (AO) | HPLC | Assay by HPLC | % RSD failure for bracketing standard solution | Finished product | During the analysis of Aripiprazole, Assay by HPLC’, it was observed that, in bracketing Standard solution-02 injection, % RSD for the areas of Aripiprazole peak not meet the acceptance criteria. Observed %RSD:1.95 %, against Acceptance criteria: NMT: 1.0%. | Based on the investigation, bracketing standard solution-02 injection %RSD for the areas of Aripiprazole peak not met the acceptance criteria due to bracketing standard solution-02 injection area observed higher side due to analyst overtightened the bracketing standard solution-02 vial by inadvertently and leads to the increased area. | Human Error
86 | Losartan Potassium (Form-I) (LR) | HPLC | Related Substances/Compounds by HPLC | weights not entered in custom fields | Finished product | While running the Losartan Potassium Related Compounds by HPLC analysis, impurity-D weight not entered in custom field amounts table. | Based on the above investigation, while running of analysis, it was observed that, analyst prepared the Reference solution (c) (i.e. containing Impurity-D) and missed to enter the impurity-D standard weight in custom fields amounts table by inadvertently and started the existing reviewed sequence. | Human Error
87 | Levetiracetam (Form-I) (LE) | Particle size analyzer | Particle Size by Laser Diffraction | % RSD failure | Finished product | During the analysis of Levetiracetam (Form-I) Particle size by Laser diffraction (Wet method) for the batch LE24080286, it is observed that, %RSD not met the acceptance criteria at d (0.9). Observed value:25.9%, Acceptance criteria NMT:15.0%. | NA | NA
88 | Lopinavir (LPR/LQ) | HPLC | Related Substances/Compounds by HPLC | % RSD failure for bracketing standard solution | Finished product | During the analysis of Lopinavir Related substances by HPLC (Procedure-1), it is noticed that in 2nd bracketing standard solution injection %RSD for Lopinavir peak not met the acceptance criteria. Observed value is 10.01 against acceptance criteria NMT 3.0%. | Based on the investigation, in initial analysis at 2nd bracketing standard solution, %RSD for the areas of Lopinavir peak not met the acceptance criteria due to septa gone bad. Hence, there might be chance to get evaporation of diluent from vial and it leads to increasing the main peak area of 2nd bracketing standard solution injection. However, analyst not ensured the vial septa condition prior to analysis. | Human Error
89 | Hydralazine Hydrochloride (HH) | Gas Chromatography | Residual solvents by GC | Instrument Error / Failure observed | Finished product | During the analysis of Hydralazine HCl Residual solvents by GC, it was identified that injection not taken for B.No: HH24080239 Prep-02 and chromatogram containing no data. | Based on the investigation, while running of analysis, injection not taken for B. No: HH24080239 Prep-02 and chromatogram containing no data due to inadvertently pressed the restart icon by Mr. N.J.S. Raju while doing ‘login’ to desktop. | Human Error
90 | Dolutegravir Sodium (DU) | HPLC | Assay by HPLC | Instrument Error / Failure observed | Finished product | During the analysis of Dolutegravir Sodium Assay by HPLC (On anhydrous basis) while running DU24080030 Pre’n-01, it is noticed that instrument failure occurred at around 10 minutes. | Based on the investigation, initial analysis at 3rd batch (DU24080030 Pre-01) instrument failure occurred might be due to network drop between instrument and computer. | Instrument Error
91 | Levetiracetam (Form-I) (LE) | Particle size analyzer | Particle Size by Laser Diffraction | % RSD failure | In-Process samples | During the analysis of Levetiracetam (Form-I) Particle size by Laser diffraction (Wet method) for the batch LE24080311-INP preparation-01, it is observed that, %RSD not met the acceptance criteria at d (0.9). Observed value:22.3%, Acceptance criteria NMT:15.0%. | NA | NA
92 | Atomoxetine Hydrochloride (AH) | Gas Chromatography | Residual solvents by GC | Instrument Error / Failure observed | In-Process samples | During the analysis of AHC-II Residual solvents by GC, it was identified that while running the B.No: AHC2240032 ( 3rd interval TOP sample) instrument failure occurred at around 22 minutes, as per STP total run time is 25 minutes. | NA | NA